# FRDM-MCXA153 4-20 mA Handoff

Oct 26, 2026 · @dom

The FRDM-MCXA153 drives a calibrated 4-20 mA loop through a NOYITO V/I module on a 20 V supply; calibration is saved in flash, with about ±0.15 mA error mid-scale.

## Quick start

The firmware and calibration are already on the board; nothing needs building or flashing.

- Connect a USB cable from the computer to the board's J15 USB-C port (labelled MCU-Link).

- Open the serial port: picocom -b 115200 /dev/ttyACM0 (exit with Ctrl-A, then Ctrl-X). Press Enter to get the > prompt.

- Set the loop current with ma, e.g. ma 12.5. ma alone shows the present setting.

> ma 12.5
ma: 12.499 mA (532/1000 counts)
> ma
ma: 12.499 mA (532/1000 counts)

The 20 V loop supply and wiring below must be in place for current to flow. Usable range is about 3.85 to 20.35 mA; values outside 0 to 25 mA are rejected. If /dev/ttyACM0 is denied, add the user to the plugdev group (or dialout, depending on the system) and log in again.

## Board state

The board runs the blinky + serial shell firmware from ~/proj/MCXA153/blinky on the Pi, and the calibration below is saved in flash.

| Item | State |
| --- | --- |
| Firmware | build/blinky.elf, built 2026-10-22, running (core 48 MHz) |
| Saved settings | 1 of 170 flash slots used, written 2026-10-26; restored at power-up |
| Output when left | ma 4 = 26 counts, about 4.0 mA |
| LED | Cycle mode, 500 ms (default) |
| Loop supply | 20 V bench supply |
| Source code | Plain folder on the Pi, not under git |

## Wiring

The PWM on J1-2 is filtered by 10 kΩ + 2.2 µF into the module's VIN; the module sources the loop from the 20 V supply. The 20 V never touches the FRDM board; only grounds are shared.

| From | To | Notes |
| --- | --- | --- |
| FRDM J1-2 (P1_4, PWM) | 10 kΩ → module VIN+ | 12 kHz, 1000 counts, 0–3.3 V |
| Module VIN+ | 2.2 µF electrolytic → GND | − leg to GND |
| Module VIN− | FRDM GND (J6-8) | Third input terminal not connected |
| Supply + (20 V) | Module 24V terminal | Module accepts 12–24 V |
| Supply − | Module 0V terminal and FRDM GND (J6-8) | Common ground |
| Module AOUT+ (silkscreen AO+) | mA meter → 250 Ω load → module 0V | 1 V at 4 mA, 5 V at 20 mA across the load |
| FRDM USB-C J15 (MCU-Link) | Raspberry Pi USB | Power, flashing, serial /dev/ttyACM0 |

Measured VIN at 50% PWM with the supply off: 1.60 V (ideal 1.65 V; the 50 mV loss is input loading through the 10 kΩ and is absorbed by calibration).

## Calibration results

The module pots were trimmed to about 3.85–20.5 mA, then a two-point calibration was saved: 25 counts = 3.990 mA, 970 counts = 19.850 mA (16.8 µA per count).

| PWM counts | Purpose | Firmware predicts (mA) | Meter read (mA) |
| --- | --- | --- | --- |
| 0 | Zero trim | — | 3.80, then 3.85 after span trim |
| 10 | Low-end probe | — | 3.88 |
| 20 | Low-end probe | — | 3.96 |
| 25 | Cal lo point | 3.990 | 3.99 |
| 502 | Mid check (ma 12) | 11.996 | 11.85 |
| 970 | Cal hi point | 19.850 | 19.85 |
| 1000 | Span trim | — | 20.50 |

Below about 20 counts the module barely responds (+0.03 mA over the first 10 counts against about 0.17 mA expected), so the low point was taken at 25 counts. The firmware's "reachable 3.57 mA" figure is a straight-line extrapolation; the real floor is about 3.85 mA.

## Operating it

Open the shell with picocom -b 115200 /dev/ttyACM0 (exit Ctrl-A, Ctrl-X); only one program can hold the port. Build and flash from ~/proj/MCXA153/blinky with make && make flash; flashing keeps the saved calibration. Full details are in that folder's README.md.

| Command | Effect |
| --- | --- |
| ma <mA> | Set loop current, e.g. ma 12.5 (0.001 mA steps, about 17 µA resolution) |
| ma | Show predicted loop current |
| pwm <0-100> / pwm counts <n> | Raw PWM duty, bypasses calibration |
| cal | Show calibration and reachable range |
| cal lo <mA> / cal hi <mA> | Record the meter reading at the present output (near 4 / 20 mA) |
| cal reset | Back to 0 counts = 4 mA, 1000 counts = 20 mA |
| save | Store settings and calibration in flash |
| defaults | Erase saved settings, including calibration |
| status | Mode, uptime, PWM and mA |

To recalibrate: pwm counts 25, read meter, cal lo <reading>; pwm counts 970, read meter, cal hi <reading>; check ma 12; save.

## Sample commands and output

This output was captured from the board on 2026-10-26 with the saved calibration. The firmware prints a predicted current; the meter column shows what the loop actually carries.

| Command | Counts set | Firmware prints | Meter (mA) |
| --- | --- | --- | --- |
| ma 4 | 26 | 4.007 mA | 4.01 measured |
| ma 12 | 502 | 11.996 mA | 11.85 measured |
| ma 20 | 979 | 20.001 mA | 19.92 measured (0.08 low) |
| pwm 50 | 500 | 1.650 V avg | about 11.82 expected |
| pwm counts 970 | 970 | 3.201 V avg | 19.85 measured |
| ma 30 | unchanged | usage error | unchanged |

Full console session (picocom, 115200 8N1):

> status
mode=cycle rate=500ms uptime=2218s core=48000000Hz rx_overruns=0 rx_errors=0
pwm: 26/1000 counts (2.6%) @ 12000 Hz -> 0.086 V avg on P1_4 (J1-2)
ma: 4.007 mA (26/1000 counts)
> cal
cal: 25 counts = 3.990 mA, 970 counts = 19.850 mA
     reachable 3.571 mA .. 20.353 mA, 16783 nA/count
> ma 4
ma: 4.007 mA (26/1000 counts)
> ma 12
ma: 11.996 mA (502/1000 counts)
> ma 20
ma: 20.001 mA (979/1000 counts)
> ma
ma: 20.001 mA (979/1000 counts)
> ma 30
err: usage: ma <0.000..25.000>
> pwm 50
pwm: 500/1000 counts (50.0%) @ 12000 Hz -> 1.650 V avg on P1_4 (J1-2)
> pwm counts 970
pwm: 970/1000 counts (97.0%) @ 12000 Hz -> 3.201 V avg on P1_4 (J1-2)
> pwm
pwm: 970/1000 counts (97.0%) @ 12000 Hz -> 3.201 V avg on P1_4 (J1-2)

The V avg figure assumes 3.300 V I/O and no load; at the module's VIN expect about 3% less (1.60 V measured at 50%). ma with no argument and pwm with no argument only report; they change nothing.

## Open items

The main limit is a 0.15 mA sag at mid-scale (about 0.9% of span) that a two-point calibration cannot remove.

- ☒ Read the meter at ma 4 (26 counts): 4.01 mA measured

- ☒ Run ma 20 (979 counts): 19.92 mA measured, 0.08 mA below prediction

- ☐ Map the error at 8 and 16 mA to see the shape of the sag

- ☐ Decide whether ±0.15 mA is acceptable; if not, add a multi-point (piecewise-linear) calibration in main.c and settings.c (changes the flash record format, so recalibrate after)

- ☐ Put ~/proj/MCXA153/blinky under git before further changes